## Description
The BathRC model: a method to estimate flow restrictor size for dual ventilation of dissimilar patients
A R Plummer, University of Bath, 2nd April 2020
v4 14/4/2020
For dual patient ventilation (also known as ventilator splitting or sharing), we need to add a flow restrictor in the inspiration tube for the patient who is more compliant, or requires a lower tidal volume. We have shown that a simple linear resistance-compliance model, termed the BathRC model, of the ventilator tubing system and lung allows direct calculation of the flow restriction resistance from ventilator pressures and the required tidal volume. The BathRC model has been experimentally validated using a GE Aisys CS2 ventilator connected to two test lungs.
Dual patient restrictor calc sheet. A patient can be characterised by a series resistance (R) and compliance (C). Ventilator tubing system can also be approximated by lumped linear model, i.e resistance (Rv) and compliance (Cv). When two patients are hooked up to the same ventilator, pressure control will be used, the ventilator stepping between low pressure (PEEP) and high pressure (PEEP+Pinsp). So the tidal volume (Vmax-Vmin) delivered to each patient, from minimum lung volume during expiration (Vmin), to maximum lung volume during inspiration (Vmax), can no longer be individually controlled by the ventilator pressure settings alone. This sheet calculates the volumes for each patient (A and B). Also, for a given desired tidal volume for patient A, the value of an additional flow restrictor resistance in the patient's inspiration line is calculated. Thus ventilator pressure rise (Pinsp) should be set to give required tidal volume in patient B, and then the restrictor resistance to reduce the tidal volume in patient A is calculated. So pateient A must be the more compliant, or require the lower tidal volume. NOTE : lung volume is taken as zero at zero gauge pressure. FRC is not included.
EXAMPLE. An example ventilator pressure control setting might be described as 30/10 cmH20, meaning Pinsp=20 cmH20, and PEEP = 10 cmH20 (so ventilator switches between 30cmH20 and 10cmH20, because Pinsp is the pressure rise above PEEP). Then you get different tidal volume ranges (row 32) for the 2 patients with these pressure settings, dependent on patient lung compliance and airway resistance. Adding an inspiration restrictor can only reduce the tidal volume, not increase it. Putting the same value as the calculated tidal volume in the entry field for tidal volume (row 35) will give a restrictor resistance of zero, or a smaller value of tidal volume gives a sensible restriction. Putting a larger tidal volume gives a negative restriction (as its trying to supercharge the patient to get more volume in! ... or in the extreme it can't calculate a value at all).
Multi-patient comparison sheet. For 'advanced users'. Allows a restrictor resistance value to be specified a priori if desired. Provides some additional calulations too.

## Dual patient restrictor calc
BathRC model
Calculation of inspiration restrictor resistance to alter patient A ventilation characteristics. Patient B is unaltered.
Of use for ventilator sharing between patients |  | White fields are for data entry |  |  |  |  | Values for plot
[6 empty] | Time (s) | A, no restr. | B | A, with restr.
Ventilator settings. Ventilator must be in Pressure Control Ventilation (PCV) mode |  |  |  | Derived ventilator parameters |  | 0 | 0.46816 | 0.14619 | 0.42331
Respiration rate (breaths / min) | 15 |  |  | Respiration period (s) | 4 | 0.05 | 0.49691 | 0.17629 | 0.44353
Inspiration to expiration time ratio (I/E) | 0.5 |  |  | I to E switching time within period (s) | 1.3333 | 0.1 | 0.52495 | 0.20481 | 0.46341
Pinsp (inspiration pressure above PEEP, cmH2O) | 25 |  |  | Max ventilator pressure (cmH2O) | 30 | 0.15 | 0.55228 | 0.23183 | 0.48295
PEEP (cmH2O) | 5 |  |  |  |  | 0.2 | 0.57893 | 0.25744 | 0.50216
Tubing system resistance (cmH2O/(L/s) ) | 22 |  |  |  |  | 0.25 | 0.60492 | 0.2817 | 0.52105
Tubing system compliance (L/cmH2O) | 0.004 |  |  |  |  | 0.3 | 0.63026 | 0.30469 | 0.53962
[6 empty] | 0.35 | 0.65496 | 0.32648 | 0.55787
Patient characteristics | A | B |  |  |  | 0.4 | 0.67905 | 0.34712 | 0.57582
Inspiration resistance, Ri (cmH2O/(L/s) ) | 13 | 13 | Note: 13 is median for normal lung, Arnal 2018 |  |  | 0.45 | 0.70254 | 0.36668 | 0.59346
Expiration resistance, Re (cmH2O/(L/s) ) | 13 | 13 | Note: same as inspiration often used |  |  | 0.5 | 0.72544 | 0.38521 | 0.6108
Inspiration Compliance, Ci (L/cmH2O) | 0.054 | 0.024 | Note: 0.054 is median for normal lung, Arnal 2018 |  |  | 0.55 | 0.74777 | 0.40277 | 0.62786
Expiration Compliance, Ce (L/cmH2O) | 0.054 | 0.024 | Note: same as inspiration often used |  |  | 0.6 | 0.76954 | 0.41941 | 0.64462
Derived patient parameters | [5 empty] | 0.65 | 0.79077 | 0.43518 | 0.6611
Patient inspiration time constant, RiCi (s) | 0.702 | 0.312 |  |  |  | 0.7 | 0.81147 | 0.45012 | 0.6773
Patient expiration time constant, ReCe (s) | 0.702 | 0.312 |  |  |  | 0.75 | 0.83165 | 0.46427 | 0.69323
Derived system parameters | [5 empty] | 0.8 | 0.85133 | 0.47769 | 0.70889
Ventilation inspiration time constant (s) | 1.978 | 0.928 |  |  |  | 0.85 | 0.87051 | 0.4904 | 0.72428
Ventilation expiration time constant (s) | 1.978 | 0.928 |  |  |  | 0.9 | 0.88922 | 0.50244 | 0.73942
Steady end-inspiration lung volume (L) | 1.62 | 0.72 |  |  |  | 0.95 | 0.90746 | 0.51385 | 0.75429
Steady end-expiration lung volume (L) | 0.27 | 0.12 |  |  |  | 1 | 0.92525 | 0.52467 | 0.76892
Coefficient a | 0.02648 | 0.018295 |  |  |  | 1.05 | 0.94259 | 0.53491 | 0.7833
Coefficient b | 0.50962 | 0.23769 |  |  |  | 1.1 | 0.9595 | 0.54462 | 0.79744
Coefficient c | 0.039975 | 0.022644 |  |  |  | 1.15 | 0.97599 | 0.55382 | 0.81133
Coefficient d | 0.25972 | 0.056497 |  |  |  | 1.2 | 0.99206 | 0.56254 | 0.825
Max lung volume, in addition to FRC (L) | 1.033 | 0.58361 |  |  |  | 1.25 | 1.0077 | 0.5708 | 0.83843
Min lung volume, in addition to FRC (L) | 0.46816 | 0.14619 |  |  |  | 1.3 | 1.023 | 0.57862 | 0.85164
Tidal volume (L) | 0.56483 | 0.43742 |  |  |  | 1.35 | 1.0266 | 0.57536 | 0.85536
[6 empty] | 1.4 | 1.0077 | 0.55147 | 0.84075
Inspiration restriction to achieve a specified tidal volume | [5 empty] | 1.45 | 0.98929 | 0.52884 | 0.8265
Required tidal volume (L). Must be <= value above | 0.437 |  |  |  |  | 1.5 | 0.97134 | 0.5074 | 0.81261
New max lung volume, in addition to FRC (L) | 0.86031 |  |  |  |  | 1.55 | 0.95383 | 0.48708 | 0.79907
New min lung volume, in addition to FRC (L) | 0.42331 |  |  |  |  | 1.6 | 0.93676 | 0.46782 | 0.78586
New coefficient b | 0.63482 |  |  |  |  | 1.65 | 0.92012 | 0.44958 | 0.77299
New ventilation inspiration time constant (s) | 2.9342 |  |  |  |  | 1.7 | 0.90389 | 0.43229 | 0.76043
Inspiration restrictor needed (cmH2O/(L/s)) | 16.4867 |  |  |  |  | 1.75 | 0.88807 | 0.41591 | 0.74819
[6 empty] | 1.8 | 0.87264 | 0.40039 | 0.73625
[6 empty] | 1.85 | 0.8576 | 0.38568 | 0.72461
[6 empty] | 1.9 | 0.84293 | 0.37174 | 0.71327
[6 empty] | 1.95 | 0.82863 | 0.35854 | 0.7022
[6 empty] | 2 | 0.81469 | 0.34603 | 0.69141
[6 empty] | 2.05 | 0.80109 | 0.33417 | 0.68089
[6 empty] | 2.1 | 0.78783 | 0.32294 | 0.67064
[6 empty] | 2.15 | 0.77491 | 0.31229 | 0.66064
[6 empty] | 2.2 | 0.76231 | 0.30221 | 0.65089
[6 empty] | 2.25 | 0.75002 | 0.29265 | 0.64138
[6 empty] | 2.3 | 0.73804 | 0.28359 | 0.63211
[6 empty] | 2.35 | 0.72635 | 0.27501 | 0.62307
[6 empty] | 2.4 | 0.71496 | 0.26688 | 0.61426
[6 empty] | 2.45 | 0.70385 | 0.25918 | 0.60566
[6 empty] | 2.5 | 0.69302 | 0.25188 | 0.59729
[6 empty] | 2.55 | 0.68247 | 0.24496 | 0.58912
[6 empty] | 2.6 | 0.67217 | 0.2384 | 0.58115
[6 empty] | 2.65 | 0.66213 | 0.23219 | 0.57338
[6 empty] | 2.7 | 0.65234 | 0.22631 | 0.56581
[6 empty] | 2.75 | 0.6428 | 0.22073 | 0.55843
[6 empty] | 2.8 | 0.63349 | 0.21545 | 0.55123
[6 empty] | 2.85 | 0.62442 | 0.21044 | 0.54421
[6 empty] | 2.9 | 0.61557 | 0.2057 | 0.53736
[6 empty] | 2.95 | 0.60695 | 0.2012 | 0.53069
[6 empty] | 3 | 0.59854 | 0.19694 | 0.52418
[6 empty] | 3.05 | 0.59034 | 0.19291 | 0.51784
[6 empty] | 3.1 | 0.58234 | 0.18908 | 0.51165
[6 empty] | 3.15 | 0.57454 | 0.18546 | 0.50562
[6 empty] | 3.2 | 0.56694 | 0.18203 | 0.49974
[6 empty] | 3.25 | 0.55953 | 0.17877 | 0.494
[6 empty] | 3.3 | 0.5523 | 0.17569 | 0.48841
[6 empty] | 3.35 | 0.54526 | 0.17277 | 0.48296
[6 empty] | 3.4 | 0.53839 | 0.17 | 0.47764
[6 empty] | 3.45 | 0.53169 | 0.16738 | 0.47246
[6 empty] | 3.5 | 0.52515 | 0.16489 | 0.46741
[6 empty] | 3.55 | 0.51879 | 0.16254 | 0.46248
[6 empty] | 3.6 | 0.51258 | 0.16031 | 0.45768
[6 empty] | 3.65 | 0.50652 | 0.15819 | 0.45299
[6 empty] | 3.7 | 0.50062 | 0.15619 | 0.44842
[6 empty] | 3.75 | 0.49486 | 0.15429 | 0.44397
[6 empty] | 3.8 | 0.48925 | 0.15249 | 0.43963
[6 empty] | 3.85 | 0.48377 | 0.15079 | 0.43539
[6 empty] | 3.9 | 0.47844 | 0.14917 | 0.43126
[6 empty] | 3.95 | 0.47324 | 0.14764 | 0.42724
[6 empty] | 4 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.05 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.1 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.15 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.2 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.25 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.3 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.35 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.4 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.45 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.5 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.55 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.6 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.65 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.7 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.75 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.8 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.85 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.9 | 0.46816 | 0.14619 | 0.42331
[6 empty] | 4.95 | 0.46816 | 0.14619 | 0.42331

## Multi-patient comparison
BathRC model
Calculation of inspiration restrictor resistance to alter patient ventilation characteristics. Comparison of 4 patients
Includes extra calculations. Advanced users only! |  | White fields are for data entry
[7 empty] | Time (s) | 1 | 2 | 3 | 4
Ventilator settings. Ventilator must be in Pressure Control Ventilation (PCV) mode |  |  |  |  | Derived ventilator parameters |  | 0 | 0.46816 | 0.35075 | 0.24145 | 0.14619
Respiration rate (breaths / min) | 15 |  |  |  | Respiration period (s) | 4 | 0.05 | 0.49691 | 0.38006 | 0.27132 | 0.17629
Inspiration to expiration time ratio (I/E) | 0.5 |  |  |  | I to E switching time within period (s) | 1.3333 | 0.1 | 0.52495 | 0.40849 | 0.30005 | 0.20481
Pinsp (inspiration pressure above PEEP, cmH2O) | 25 |  |  |  | Max ventilator pressure (cmH2O) | 30 | 0.15 | 0.55228 | 0.43606 | 0.32767 | 0.23183
PEEP (cmH2O) | 5 | [5 empty] | 0.2 | 0.57893 | 0.4628 | 0.35424 | 0.25744
Tubing system resistance (cmH2O/(L/s) ) | 22 | [5 empty] | 0.25 | 0.60492 | 0.48872 | 0.37978 | 0.2817
Inspiration flow restrictor (cmH2O/(L/s) ) | 0 | This is an additional flow restriction which affects all 4 patients |  |  |  |  | 0.3 | 0.63026 | 0.51387 | 0.40434 | 0.30469
Tubing system compliance (L/cmH2O) | 0.004 | [5 empty] | 0.35 | 0.65496 | 0.53825 | 0.42797 | 0.32648
[7 empty] | 0.4 | 0.67905 | 0.56189 | 0.45068 | 0.34712
Patient characteristics | A | B | C | D |  |  | 0.45 | 0.70254 | 0.58482 | 0.47252 | 0.36668
Inspiration resistance, Ri (cmH2O/(L/s) ) | 13 | 13 | 13 | 13 | Note: 13 is median for normal lung, Arnal 2018 |  | 0.5 | 0.72544 | 0.60706 | 0.49353 | 0.38521
Expiration resistance, Re (cmH2O/(L/s) ) | 13 | 13 | 13 | 13 | Note: same as inspiration often used |  | 0.55 | 0.74777 | 0.62862 | 0.51373 | 0.40277
Inspiration Compliance, Ci (L/cmH2O) | 0.054 | 0.044 | 0.034 | 0.024 | Note: 0.054 is median for normal lung, Arnal 2018 |  | 0.6 | 0.76954 | 0.64953 | 0.53315 | 0.41941
Expiration Compliance, Ce (L/cmH2O) | 0.054 | 0.044 | 0.034 | 0.024 | Note: same as inspiration often used |  | 0.65 | 0.79077 | 0.66981 | 0.55183 | 0.43518
Derived patient parameters | [6 empty] | 0.7 | 0.81147 | 0.68948 | 0.5698 | 0.45012
Patient inspiration time constant, RiCi (s) | 0.702 | 0.572 | 0.442 | 0.312 |  |  | 0.75 | 0.83165 | 0.70855 | 0.58707 | 0.46427
Patient expiration time constant, ReCe (s) | 0.702 | 0.572 | 0.442 | 0.312 |  |  | 0.8 | 0.85133 | 0.72704 | 0.60368 | 0.47769
Derived system parameters | [6 empty] | 0.85 | 0.87051 | 0.74498 | 0.61965 | 0.4904
Ventilation inspiration time constant (s) | 1.978 | 1.628 | 1.278 | 0.928 |  |  | 0.9 | 0.88922 | 0.76237 | 0.63501 | 0.50244
Ventilation expiration time constant (s) | 1.978 | 1.628 | 1.278 | 0.928 |  |  | 0.95 | 0.90746 | 0.77923 | 0.64979 | 0.51385
Steady end-inspiration lung volume (L) | 1.62 | 1.32 | 1.02 | 0.72 |  |  | 1 | 0.92525 | 0.79559 | 0.66399 | 0.52467
Steady end-expiration lung volume (L) | 0.27 | 0.22 | 0.17 | 0.12 |  |  | 1.05 | 0.94259 | 0.81145 | 0.67765 | 0.53491
Coefficient a | 0.02648 | 0.024602 | 0.022022 | 0.018295 |  |  | 1.1 | 0.9595 | 0.82683 | 0.69079 | 0.54462
Coefficient b | 0.50962 | 0.44087 | 0.35229 | 0.23769 |  |  | 1.15 | 0.97599 | 0.84175 | 0.70342 | 0.55382
Coefficient c | 0.039975 | 0.035448 | 0.02978 | 0.022644 |  |  | 1.2 | 0.99206 | 0.85621 | 0.71556 | 0.56254
Coefficient d | 0.25972 | 0.19437 | 0.12411 | 0.056497 |  |  | 1.25 | 1.0077 | 0.87024 | 0.72724 | 0.5708
Max lung volume, in addition to FRC (L) | 1.033 | 0.89268 | 0.74572 | 0.58361 |  |  | 1.3 | 1.023 | 0.88384 | 0.73848 | 0.57862
Min lung volume, in addition to FRC (L) | 0.46816 | 0.35075 | 0.24145 | 0.14619 |  |  | 1.35 | 1.0266 | 0.88583 | 0.73827 | 0.57536
Tidal volume (L) | 0.56483 | 0.54194 | 0.50427 | 0.43742 |  |  | 1.4 | 1.0077 | 0.86569 | 0.71646 | 0.55147
[7 empty] | 1.45 | 0.98929 | 0.84616 | 0.6955 | 0.52884
Inspiration restriction to achieve a specified tidal volume | [6 empty] | 1.5 | 0.97134 | 0.82723 | 0.67533 | 0.5074
Required tidal volume (L). Must be <= value above | 0.437 | 0.437 | 0.437 | 0.437 |  |  | 1.55 | 0.95383 | 0.80886 | 0.65594 | 0.48708
New max lung volume, in addition to FRC (L) | 0.86031 | 0.76243 | 0.66892 | 0.58317 |  |  | 1.6 | 0.93676 | 0.79105 | 0.6373 | 0.46782
New min lung volume, in addition to FRC (L) | 0.42331 | 0.32543 | 0.23192 | 0.14617 |  |  | 1.65 | 0.92012 | 0.77378 | 0.61937 | 0.44958
New coefficient b | 0.63482 | 0.56061 | 0.44549 | 0.23845 |  |  | 1.7 | 0.90389 | 0.75703 | 0.60213 | 0.43229
New ventilation inspiration time constant (s) | 2.9342 | 2.3039 | 1.649 | 0.93007 |  |  | 1.75 | 0.88807 | 0.74079 | 0.58555 | 0.41591
Inspiration restrictor needed (cmH2O/(L/s)) | 16.4867 | 14.0817 | 9.7623 | 0.073992 |  |  | 1.8 | 0.87264 | 0.72503 | 0.5696 | 0.40039
OR inspiration added compliance needed (L/cmH2O) | 0.043465 | 0.030724 | 0.016862 | 9.4171e-05 |  |  | 1.85 | 0.8576 | 0.70976 | 0.55427 | 0.38568
[7 empty] | 1.9 | 0.84293 | 0.69495 | 0.53953 | 0.37174
Inspiration restriction to achieve the same tidal volume (row 33) with a higher Pinsp pressure | [6 empty] | 1.95 | 0.82863 | 0.68058 | 0.52535 | 0.35854
Higher Pinsp (insp. pressure above PEEP, cmH2O) | 30 | 30 | 30 | 30 |  |  | 2 | 0.81469 | 0.66665 | 0.51172 | 0.34603
New steady end-inspiration lung volume (L) | 1.89 | 1.54 | 1.19 | 0.84 |  |  | 2.05 | 0.80109 | 0.65314 | 0.49861 | 0.33417
New coefficient b | 0.60274 | 0.54431 | 0.46837 | 0.36954 |  |  | 2.1 | 0.78783 | 0.64004 | 0.486 | 0.32294
New ventilation inspiration time constant (s) | 2.6337 | 2.1921 | 1.7579 | 1.3394 |  |  | 2.15 | 0.77491 | 0.62734 | 0.47387 | 0.31229
Inspiration restrictor needed (cmH2O/(L/s)) | 11.3048 | 11.7521 | 12.6286 | 14.6915 |  |  | 2.2 | 0.76231 | 0.61502 | 0.46221 | 0.30221
[7 empty] | 2.25 | 0.75002 | 0.60307 | 0.451 | 0.29265
[7 empty] | 2.3 | 0.73804 | 0.59148 | 0.44022 | 0.28359
[7 empty] | 2.35 | 0.72635 | 0.58025 | 0.42985 | 0.27501
[7 empty] | 2.4 | 0.71496 | 0.56935 | 0.41988 | 0.26688
[7 empty] | 2.45 | 0.70385 | 0.55878 | 0.4103 | 0.25918
[7 empty] | 2.5 | 0.69302 | 0.54854 | 0.40108 | 0.25188
[7 empty] | 2.55 | 0.68247 | 0.5386 | 0.39221 | 0.24496
[7 empty] | 2.6 | 0.67217 | 0.52896 | 0.38368 | 0.2384
[7 empty] | 2.65 | 0.66213 | 0.51962 | 0.37549 | 0.23219
[7 empty] | 2.7 | 0.65234 | 0.51056 | 0.3676 | 0.22631
[7 empty] | 2.75 | 0.6428 | 0.50177 | 0.36002 | 0.22073
[7 empty] | 2.8 | 0.63349 | 0.49325 | 0.35273 | 0.21545
[7 empty] | 2.85 | 0.62442 | 0.48498 | 0.34572 | 0.21044
[7 empty] | 2.9 | 0.61557 | 0.47697 | 0.33898 | 0.2057
[7 empty] | 2.95 | 0.60695 | 0.4692 | 0.33249 | 0.2012
[7 empty] | 3 | 0.59854 | 0.46166 | 0.32626 | 0.19694
[7 empty] | 3.05 | 0.59034 | 0.45435 | 0.32026 | 0.19291
[7 empty] | 3.1 | 0.58234 | 0.44726 | 0.3145 | 0.18908
[7 empty] | 3.15 | 0.57454 | 0.44039 | 0.30895 | 0.18546
[7 empty] | 3.2 | 0.56694 | 0.43372 | 0.30362 | 0.18203
[7 empty] | 3.25 | 0.55953 | 0.42726 | 0.2985 | 0.17877
[7 empty] | 3.3 | 0.5523 | 0.42099 | 0.29356 | 0.17569
[7 empty] | 3.35 | 0.54526 | 0.41491 | 0.28882 | 0.17277
[7 empty] | 3.4 | 0.53839 | 0.40901 | 0.28426 | 0.17
[7 empty] | 3.45 | 0.53169 | 0.4033 | 0.27988 | 0.16738
[7 empty] | 3.5 | 0.52515 | 0.39775 | 0.27566 | 0.16489
[7 empty] | 3.55 | 0.51879 | 0.39238 | 0.27161 | 0.16254
[7 empty] | 3.6 | 0.51258 | 0.38716 | 0.26771 | 0.16031
[7 empty] | 3.65 | 0.50652 | 0.38211 | 0.26396 | 0.15819
[7 empty] | 3.7 | 0.50062 | 0.3772 | 0.26036 | 0.15619
[7 empty] | 3.75 | 0.49486 | 0.37245 | 0.25689 | 0.15429
[7 empty] | 3.8 | 0.48925 | 0.36784 | 0.25356 | 0.15249
[7 empty] | 3.85 | 0.48377 | 0.36337 | 0.25035 | 0.15079
[7 empty] | 3.9 | 0.47844 | 0.35903 | 0.24727 | 0.14917
[7 empty] | 3.95 | 0.47324 | 0.35483 | 0.2443 | 0.14764
[7 empty] | 4 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.05 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.1 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.15 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.2 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.25 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.3 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.35 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.4 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.45 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.5 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.55 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.6 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.65 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.7 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.75 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.8 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.85 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.9 | 0.46816 | 0.35075 | 0.24145 | 0.14619
[7 empty] | 4.95 | 0.46816 | 0.35075 | 0.24145 | 0.14619
Note on A42: Andy: Alternative modification to get vol. of Row 35. Caveat: to affect inspiration only, needs to be distal side of exp. check valve, resistance of which has not been accounted for.